## Notes
Surgery Notes
Date | Surgery | Notes | Problems with Instruments | Solutions
2016-09-15 | Cholesteatoma removal - repeat surgery | - as per scan, cholesteatoma was in the mastoid bone which needs to be accessed via incision, so not primarily an endoscopic case - even though used incision, endoscope was useful and used mostly to visualize inside the ear and especially into the middle ear - incision allows wider acecss, and better visibility of the cholesteatoma | - using a thin metal tip (looked like an unbent hockey stick) - took a few moves and a few changes of instruments to grip the specimen as it was slipping - need better grip on the instrument - instruments could not reach around the tip of the endoscope to access portions visualized by endoscope - even instruments that were rigid and curved or angled (double or single angle) couldn't reach some places - perhaps a flexible or bend-controlling instrument will be better - forceps - one has a short tip and cannot go around endoscope, other has longer tip and has trouble reaching around endoscope - Panetti dissector kept blocking with bone chips | - instrument in many sizes, angles that are optimized to reach hard to reach areas (optimize using 3D model) - use the double bend in the forceps with long tips to reach around the endoscope - Kyle's manual tool but different handle -> naturally tip unbent, with a roller that allows the tip to bend just by using one finger, roller needs to have enough friction to maintain the position but not be too tight - 'stem' can be bent, or assume different geometries to get around the endoscope and then the business end can be flexible/controlled
- concentric tubes - pre-shape of the curve and push it out when ready - ask Vivek what the sharpest curved tool he has made - show Dr. James a video on concentric tubes to get the idea - wrist needs cable inside so it will be larger diameter
2016-09-20 | Tympanoplasty - Narrow ear canal | - fellow was being trained - may be uncomfortable/unnatural to hold the anaesthesia syringe ready to be injected and inserting into ear canal | - switched out instruments with cleaning edges of perforation step many times: panetti curved suction, blade, rosen, forceps, angled forceps, | - rosen needle has nice curvature, if it could slice this would be easier - sharp edge rosen needle - curved panetti suction - what angle of curvature?? - determine using workspace study, and what is the angle between ear canal and umbo?
noisy suction pedal - need to find a quote for a less noisy one

## Time Flow Analysis
Surgery:
Tympanoplasty | Step | Date/notes | Time (min) | Date/notes | Time (min) | Date/notes | Time (min)
Cleaning Out Ear canal
Injecting Anaesthesia
Hair Trimming
Cleaning Edges of Perforation
Making Skin Incision
Raising Flap
Preparing Graft
Placing Graft
Replacing Flap
Packing Ear Canal
Surgery:
Cholesteatoma Removal | Step | Date/notes | Time (min) | Date/notes | Time (min) | Date/notes | Time (min)
Cleaning Out Ear canal
Injecting Anaesthesia
Hair Trimming
Cleaning Edges of Perforation
Making Skin Incision
Raising Flap
Preparing Graft
Placing Graft
Replacing Flap
Packing Ear Canal